APP Correlator Upgrades PAI Test Report

Version: A

Status:  Draft

2014-11-25

| Prepared By: |  |  |
| --- | --- | --- |
| Name and Signature | Organization | Signature and Date |
| Rich Lacasse Alejandro Saez Joe Greenberg | NRAO JAO NRAO |  |
| Approved By: |  |  |
| Names and Signatures | Organization | Date |
| Released By: |  |  |
| Name and Signature | Organization | Date |

Change Record

| Version | Date | Affected Section(s) | Change Request # | Reason/Initiation/Remarks |
| --- | --- | --- | --- | --- |
| A01 | 2013-10-3 |  | - | Initial draft |
| A02 | 2014-11-25 |  | - | Morph from procedure to report. Change document number. Add measured results. |

TABLE OF CONTENTS

1	Related Documents and Drawings	4

1.1	Applicable documents	4

1.2	Reference documents	4

1.3	Acronyms	4

ICD	Interface Control Document	5

2	Introduction	6

2.1	Purpose and Scope	6

2.2	Integration and Test Plan Overview	6

2.3	Detailed Test Procedures and Results	6

2.3.1	Overview	6

2.3.2	Assumptions	9

2.3.3	Special Considerations due to Using a Two Antenna Correlator	9

2.3.4	One Pulse Per Second and Clock Distribution Tests	10

2.3.4.1	Overview	10

2.3.4.2	Test of the 1-PPS Time-Difference-Monitor sub-system	10

2.3.5	Simulated Maser	12

2.3.5.1	General	13

2.3.5.2	Testing	13

2.3.6	GPS Receiver	14

2.3.6.1.1	Testing the Maser and GPS 1-PPS Signal Duty Cycle	14

2.3.7	Interface with Computing	16

2.3.7.1	Operating Temperature Check of the ROACH FPGA and PPC	17

2.3.8	Firmware and FPGA Personality Verification	19

2.3.9	Connectivity Robustness Tests	22

2.3.9.1	Correlator Card to Correlator Interface Card	22

2.3.9.2	Correlator Card to PIC	24

2.3.9.3	Ethernet to PICs	24

2.3.9.4	TE from QCC to the PIC	24

2.3.9.5	Clock Cables for PIC.	25

2.3.10	Environmental Tests	25

2.3.10.1	PIC Power Dissipation	25

2.3.10.2	1-PPS Power Dissipation	28

2.3.11	Component Burn-in Tests	28

2.3.11.1	ATX Power Supply Burn-in Tests	28

2.3.11.2	PICA Burn-in Tests	30

# Related Documents and Drawings

## Applicable documents

The following documents are part of this document to the extent specified herein. If not explicitly stated otherwise, the latest issue of the document is valid.

| AD01 | Integration and Test Plan for APP |
| --- | --- |
| AD02 | APP Project Plan |
| AD03 | APP Requirements Document |
| AD04 | APP Computing Management Plan (ALMA-05.11.60.01-001-A-PLA) |
| AD05 | Commissioning and Science Verification Plan (ALMA-05.11.10..01-001-A-PLA) |
| AD06 | Back End ICD ALMA-05.11.10.00-50.00.00.00-A-ICD |
| AD07 | Site ICD ALMA-05.11.10.00-20.00.00.00-A-ICD |
| AD08 | Correlator ICD ALMA-05.11.10.00-60.00.00.00-A-ICD |
| AD09 | Computing ICD ALMA-05.11.10.00-70.00.00.00-A-ICD |
| AD10 | General Safety Design Specification (ALMA-10.08.00.00-003-B-SPE) |
| AD11 | System Technical Requirements (ALMA-80.04.00.00-005-B-SPE) |
| AD12 | System Electrical Design Requirements (ALMA-80.05.00.00-005-C-SPE) |
| AD13 | System Electromagnetic Compatibility (EMC) Requirements (ALMA-80.05.01.00-001-B-SPE) |
| AD14 | Product Assurance Requirements (ALMA-80.11.00.00-001-D-GEN) |
| AD15 | Environmental Specification (ALMA-80.05.02.00-001-B-SPE) |
| AD16 | Operations Plan (ALMA-00.00.00.00-002-A-PLA) |
| AD17 | Seismic Design Specifications for ALMA-AOS and ALMA-OSF (SYSE-80.10.00.00-002-B-REP) |
| AD18 | ALMA Interface Management Plan (ALMA-80.07.00.00-001-D-PLA) |
| AD19 | Mark6 Users Manual (draft) |
| AD20 | APP Optical Fiber Link system design (ALMA-05.11.40.01-0001-A-DSN) |
| AD21 | Phasing Interface Card Manual (ALMA-05.11.31.05-0001-A-MAN) |
| AD22 | 2013-12-13-PIC Sequence Nrs short.xls (https://deki.mpifr-bonn.mpg.de/@api/deki/files/4486/=2013-12-13-PIC_Sequence_Nrs_short.xls) |
| AD23 | 2014-05-31-NRAO APP Misc Cales.xlsx( https://deki.mpifr-bonn.mpg.de/@api/deki/files/4487/=2014-05-31-NRAO_APP_Misc_Cables.xlsx) |

## Reference documents

| Ref. | Title | Doc. Number |
| --- | --- | --- |
| RD-01 | Incoming Inspection Report for Symmetricon XL-GPS SN1331C00739 | ALMA-05.11.35.91-0001-A-REP |
| RD-02 | PIC Assembly Check-out Logs | ALMA-05.11.31.03-0001-A-REP |

## Acronyms

A list of the acronyms used in this document is given below.

6U-PWR	6-U Power Card

9U-PWR	9-U Power Card

ALMA	Atacama Large Millimeter Array

APP	ALMA Phasing Project

ATR	Acceptance Test Report

CAN	Controller Area Network

CC		Correlator Card

CCC	Correlator Control Computer

CDP	Correlator Data Processing Computer

CDR	Critical Design Review

CVS	Concurrent Version System (software revision control system)

DMC	ACA DTS-R Monitor & Control Computer

DPI	Data Port Interface

DRX	DTS Receiver Module

DTS	Data Transmission System

DVM	Digital Volt Meter

FA		Final Adder

FPGA	Field Programmable Gate Array

GUI	Graphical User Interface

#### ICD	Interface Control Document

IPT		Integrated Product Team

LTA	Long Term Accumulator

LRU	Line Replaceable Unit

LVDS	Low Voltage Differential Signaling

NRAO	National Radio Astronomy Observatory

OPS	One PPS Distributor Board

PAS	Provisional Acceptance on Site

PLL	Phase Locked Loop

QCC	Quadrant Control Card

SC		Station Card

SCC	Station Control Card

TE 		Time Event (a 48 millisecond ALMA-wide timing signal)

TFB	Tunable Filter Board

TLC	TTL to LVDS Converter Board

# Introduction

## Purpose and Scope

This document describes procedures for doing hardware-based system tests of the ALMA Phasing Project (APP) Correlator Upgrades in Charlottesville and reports the results thereof.  See the Integration and Test Plan for APP for a description of test philosophy.

## Integration and Test Plan Overview

The purpose of the PAI tests is primarily to assure that delivered equipment meets specification before it is shipped to Chile.  Also, certain tests that only need to be done once are more easily carried out in the laboratory environment in Charlottesville.  The verification entails visual inspections, electrical and electronic tests and computer-based tests.  The flow of the testing will include:

- Verify a safe operating environment for the equipment (electrical, thermal, personnel)

- Electronic tests when applicable to verify signal characteristics

- Verification that correct firmware and FPGA personalities are installed

- Connectivity tests to assure reliable connections between installed modules

- Performance tests to assure that the equipment is performing according to the specifications.

A broader overview of the test plan is included in the Integration and Test Plan [AD-01].

## Detailed Test Procedures and Results

### Overview

For the PAI tests, a test bed was assembled in Charlottesville in the same laboratory that was used to test the 4 quadrants of the ALMA 64-Antenna Correlator.  See Figure 2-1 for a photo of the test set.  See Figure 2-2 for a block diagram of the test system.  It includes a 2-Antenna Correlator called Black Hole, a Mark 6 Recorder, the OFL system, a GPS unit, a simulated maser and a simulated 1-PPS source.  Black Hole provides a test environment (electrical, mechanical, computing) very similar to that of the 64-Antenna Correlator; it is a 2-Antenna Correlator specially modified to accept at PIC.  It also includes some computing infrastructure: an Engineering Port Computer, a CCC and CDP.  The elevation of course is not simulated, but expected performance can be extrapolated from Charlottesville to AOS.    Testing is accomplished by several means including COTS meters, probes and oscilloscopes, and built-in tests.  Details for individual tests are provided in the sections below.  Most of the testing was conducted by Alejandro Saez, Joe Greenberg and Rich Lacasse in late 2013.

Figure 2-1.  Photo of the test system used for most PAI testing.

Figure 2-2.  Block diagram of the Charlottesville test system

    - The Black Hole Two Antenna Correlator.

The test Correlator, Black Hole, is a modified version of the two Antenna Correlators which have been delivered to Chile.  A PIC, with a Roach  II attached, is included.  Also a One PPS Card is included; this is a modified Final Adder Card (same hardware but different FPGA personalities).  Two DTS Simulator Cards can inject simulated antenna signals into the TFBs.  These can be routed via the Station Cards, Station Interface Cards, Correlator Interface Cards, to the Correlator Cards.  The Correlator Card Generates Analog Sums which are routed to the PIC.   The ROACH II Card on the PIC attaches VDIF headers and outputs the sums on an optical fiber as 10 GB Ethernet via SFP+ transceivers.  A block diagram of Black Hole is shown in Figure 2-3.

Figure 2-3.  Block diagram of the test correlator, Black Hole.

### Assumptions

Black Hole has previously been checked out as a fully functioning Two Antenna Correlator.  This includes the connections of a CCC, CDP, and Engineering Port Computer.

### Special Considerations due to Using a Two Antenna Correlator

There are fundamental differences between the Two Antenna Correlator and the 64 Antenna Correlator in Chile.

The signal from each antenna is divided into 32 planes.  The 64-Antenna Correlator has separate Correlator Cards for each of the 32 planes.   Since we only have 4 Correlator Cards in the Two Antenna Correlator it is necessary to process multiple planes on each Correlator Card.  This requires special configurations and data routing.

The software group has indicated that much of their software will not work on a Two Antenna Correlator configuration.  Thus a full up test will not be available until we are at the 64 Antenna Correlator in Chile.  A full up test includes such features as sensing a phase drift then looping back a phase correction to the TFBs

### One Pulse Per Second and Clock Distribution Tests

#### Overview

Figure 2-4.  1-PPS distribution block diagram.

Figure 2-4 shows the 1PPS Distribution system.  An Oscillator provides a 5 MHz frequency reference and a 1 PPS signal.  The 5 MHz goes to a synthesizer that will generates the 125 MHz clock for Black Hole.

A GPS Receiver also provides a 1 PPS signal.  These two 1PPS signals go through a board to convert the signals to LVDS.  They then go to the 1PPS distribution board.  This is a Final Adder board with custom firmware.  In this case there is only one PIC to distribute to, but in the full correlator there will be 8 PICs.  The PIC also generates its own 1PPS from the incoming TE, and SETTIME CAN command.

#### Test of the 1-PPS Time-Difference-Monitor sub-system

Procedure: The goal of this test is to verify the functionality of the 1-PPS monitoring in the PIC.  From the Engineering Port computer issue the “dif” command in the PIC.  Using an oscilloscope, verify that the indicated time differences between the PIC 1-PPS and the Maser and GPS are reasonable.

Results:

The “dif” command shows the following:

PIC:64ANT:Quad0:Node38 > dif

Checking time difference between internal 1 PPS and external references

Type ESC or space at any time to escape.

Maser 1PPS = 0x0594B1DA [raw x8nSec] or 0.748523 secs

GPS 1PPS   = 0x0243CA03 [raw x8nSec] or 0.303452 secs

TE 1PPS    = 0x02000001 [raw x8nSec] or 0.000000 secs

A sanity check with an oscilloscope confirms that the results are reasonable.

Figure 2-5.  PIC to Maser time difference measured with an oscilloscope.

Figure 2-6.  PIC to  GPS time difference measured with an oscilloscope.

This result does not satisfy any requirements.  It is included as a matter of good practice.

### Simulated Maser

Figure 2-7.  Photo of simulated maser/1-PPS signal generator.

#### General

The signal generator shown in Figure 2-7 provides 5 MHz and a 1 PPS signal, similar to the H-Maser at AOS.  The 5 MHz goes to a synthesizer that generates the 125 MHz clock.

#### Testing

Procedure:  Verify the 125 MHz from the synthesizer provides at least 8dBm of power on a scope.  That is 1.6V Peak to Peak into a 50 Ohm load. Connect a second scope input to the 1PPS signal.  Verify it provides a TTL level 1PPS signal that is synchronous to the 125 MHz output from the synthesizer.

Hook the 125 MHz to the clock card.  Verify it syncs by observing the Red PLL Unlocked LED on the clock card go out.

Results:  The measured amplitude and frequency of the 125 MHz signal as well as the synchronization of 1-PPS to it is shown in Figure 2-8 below.

Figure 2-8.  Scope-shot of the synthesizer 125 MHz and the 1-PPS signals provided to Black Hole.

The indicator light on the clock card is off with the 125 MHz connected.  This test passes.  It does not satisfy any requirements.  It is included as a matter of good practice.

### GPS Receiver

A GPS receiver is included in the Correlator Upgrade as a redundant time check.  The Maser 1-PPS is also redundant.  Only the TE signal, ultimately derived from the Maser, is required to synchronize the Correlator for VLBI.  Having a additional 1-PPS signals as  sanity checks was considered prudent in the system design.

Procedure:

Test suite 1:  The purpose of this test is acceptance testing of the GPS receiver purchased for APP.  The procedure is in RD-01.

Test suite 2:  The purpose of this test is to verify the GPS signal chain through the 1-PPS Distributor and PIC.  This is already covered in Section 2.3.4.2.

Results: All tests passed.  Test results for suite 1 are included in RD-01.    Test suite 2 partially satisfies the following requirement: APP0200.

##### Testing the Maser and GPS 1-PPS Signal Duty Cycle

Procedure:  This test verifies connectivity of the Maser and GPS 1-PPS signals into the 1-PPS distributor and verifies that the duty cycle (technically the period of time during which the signal is high) matches that seen on an oscilloscope.  Connect to the OPS via the Engineering Port for quadrant 2 (“cc 35”).  Request the duty cycles of the GPS and Maser 1-PPS signals (“dutycycle 0” for GPS duty cycle and “dutycycle 1” for Maser duty cycle).  Remove the Maser 1-PPS signal from the TLC and repeat the dutycycle command.   Note that the maser duty cycle gets much larger, verifying connect connection all the way into the 1-PPS distributor.  Replace the Maser 1-PPS signal and remove the GPS signal.  Repeat the duty cycle command.  Note that the GPS duty cycle gets much larger, verifying correct connection all the way into the 1-PPS distributor.  Monitor both pulse widths with an oscilloscope and compare to the measured results.

Results:

Session with 1-PPS Distributor:

OPS:Quad0:Node35 > dutycycle 0

GPS

Duty Cycle = 0x000009C3

Duty Cycle = 0.019992 [mSec]

OPS:Quad0:Node35 > dutycycle 1

MASER

Duty Cycle = 0x0000000B

Duty Cycle = 0.000088 [mSec]

Oscilloscope measurements:

Figure 2-9.  Scope-shots showing Maser pulse width of 108.4 ns

Figure 2-10.  Scoppe-shots showing GPS pulse width of  19.96 us

The measurements from the dutycycle command and scope are very close.  The test passes.  This test partially satisfies APP0200.

### Interface with Computing

The test system includes a CCC and a CDP Computer.  Due to limitations of the two-antenna correlatorit is not possible to do a full loop back test in software.  That is, we cannot look at sums, calculate a phase offset and close the loop.  A Telcal computer would not be useful for the above reasons as well.  However, individual commands documented can

be tested.  A script exercising the protocols from the ICD with computing was executed successfully on August 6, 2014, per the email below.  Unfortunately, logs from the test were purged before being captured.

Figure 2-11.  Email documenting correct operation of the CAN protocols.

This satisfies the following requirements: COM-0010

### Operating Temperature Check of the ROACH FPGA and PPC

Procedure:

- Reset the correlator

- Verify that all modules respond with their prompts

- Connect to PIC-0 via the engineering port (“cc 38”)

- Command the ROACH board to power up  (“roachon”)

- After 30 seconds, connect to the PPC on the ROACH from via telnet (“telnet cockroachx 7147”), where “x” is replaced by the assigned board number.

- Force the PIC into a maximum power dissipation mode with the following sequence of commands:

  - startxmt 1

  - cc 36

  - summode 3

- Via the telnet connection monitor the PPC and FPGA temperature.

- Continue to monitor the temperature until it stabilizes (about 10 minutes) and record the temperatures

- Measure ambient temperature.

- Observe the results and verify that the operating temperatures of all FPGAs are less than 70 oC.

- Extrapolate to expected operating temperature at AOS

Results:

Transcript of telnet session:

lyra_rlacasse<1029> telnet cockroach4 7147                           	(note cockroach4 is the network name of the ROACH2 in this example)

Trying 10.12.97.220...

Connected to cockroach4.

Escape character is '^]'.

#version alpha-74-g1c6a441-dirty

#build-state 2012-12-16T22:51:01

?sensor-value raw.temp.ppc

#log info 1778100 raw 1\_clients\_subscribed\_to\_sensor\_raw.temp.ppc

#sensor-value 1778107 1 raw.temp.ppc nominal 38000						(note: this says that the temperature is 38 degrees)

!sensor-value ok 1

sensor-value raw.temp.fpga

^[[A

?sensor-value raw.temp.fpga

#log info 1808179 raw 1\_clients\_subscribed\_to\_sensor\_raw.temp.fpga

#sensor-value 1808187 1 raw.temp.fpga nominal 34000						(note: this says that the temperature is 34 degrees)

!sensor-value ok 1

Ambient temperature = 25oC.

The temperature rise of the FPGA is 9oC.

At AOS, air temperature is 16oC, air density is a factor of two lower than Charlottesville and air velocity is approximately the same.  Therefore the expected FPGA operating temperature at AOS is 16oC + 2*9oC = 34oC.  This is well below the maximum temperature of 85oC specified by Xilinx.  The test passes.  This partially satisfies APP requirements APP0070 and ENVI-00270-00.

### Firmware and FPGA Personality Verification

Procedure:  Using the engineering port, document the state all the firmware and FPGA personalities installed in the Correlator. Record screenshots of the Engineering Port session.

Results:  Every node in the system was checked on 2014-11-25.  Results for each node are shown below.

LTA:2Ant:Node0 > fpga

SA15= 00CC3228      SA31= 0062021E      SA47= 00DDF7A5      SA63= 005C391A

SA14= 004185C5      SA30= 0061F4B3      SA46= 00FF0000      SA62= 005BC084

SA13= 00CC30FA      SA29= 0060F1D5      SA45= 00FF0000      SA61= 005C6EF8

SA12= 00421023      SA28= 00613CFD      SA44= 00FF0000      SA60= 005BFC7A

SA11= 00CC227B      SA27= 006149AF      SA43= 00FF0000      SA59= 005C02EC

SA10= 004270CF      SA26= 00617F24      SA42= 00C6C46E      SA58= 005CB098

SA 9= 00CC1D55      SA25= 00611AED      SA41= 000943A7      SA57= 005C4C6B

SA 8= 0041AAC4      SA24= 00619A77      SA40= 00C6AA7A      SA56= 005C28AD

SA 7= 00CE268E      SA23= 0060EBA9      SA39= 0009883E      SA55= 005CB976

SA 6= 0042FB42      SA22= 00615D41      SA38= 00938DC9      SA54= 005C3BF4

SA 5= 00CC2329      SA21= 00618E0B      SA37= 00093053      SA53= 005C01B6

SA 4= 004239C2      SA20= 0061C317      SA36= 00327D32      SA52= 005C2F60

SA 3= 00CC2798      SA19= 00618DC5      SA35= 003271F5      SA51= 005C8F3C

SA 2= 0041A80E      SA18= 00614879      SA34= 00326D8A      SA50= 005C8BEF

SA 1= 0076352A      SA17= 006199D2      SA33= 003269B5      SA49= 005C3F78

SA 0= 0049638E      SA16= 00617D08      SA32= 0043EFD9      SA48= 005C413F

DataFlash Checksum Total = 1C255D5F

Application Checksum     = 09B9E235

BootLevel_1 Checksum     = 01A095D6

----------

TOTAL        = 277FD56A

CORR-INTF FPGA checksum = 010DB7ED

ASUMMID FPGA checksum   = 010E411D

ASUM2ND FPGA checksum   = 010E934A

ASUMEND FPGA checksum   = 010DC819

DATAOUT FPGA checksum   = 011121D0

ADDER2 FPGA checksum    = 010E5CEB

ADDER1 FPGA checksum    = 010DCFA6

LTA FPGA checksum       = 00BF98B8

LTA:2Ant:Node0 > cc 16

ASYNC_ID 16 connected.

SCC:TFB-II:Quad0:Node16 > fpga

SA15= 00EF2838      SA31= 00FF0000      SA47= 00118E0E      SA63= 00C6B4A4

SA14= 00CC2C90      SA30= 00FF0000      SA46= 000FA3A5      SA62= 000980AA

SA13= 00419740      SA29= 00FF0000      SA45= 00178022      SA61= 00C6AA7A

SA12= 00CC30FB      SA28= 00FF0000      SA44= 0010769B      SA60= 0009883E

SA11= 0040CD0F      SA27= 00FF0000      SA43= 000FDFCE      SA59= 00C6AE21

SA10= 00CC2E17      SA26= 00FF0000      SA42= 0015D500      SA58= 00093053

SA 9= 004286C3      SA25= 00FF0000      SA41= 000ED8CA      SA57= 00FF0000

SA 8= 00CC3AC8      SA24= 00FF0000      SA40= 000D6270      SA56= 00FF0000

SA 7= 0043CC17      SA23= 00FF0000      SA39= 000F9AD9      SA55= 00FF0000

SA 6= 00604BC4      SA22= 00FF0000      SA38= 000C8520      SA54= 00FF0000

SA 5= 00534B21      SA21= 00FED24B      SA37= 0013ADDB      SA53= 00FF0000

SA 4= 00506B50      SA20= 00FED149      SA36= 00144356      SA52= 00FF0000

SA 3= 0076D981      SA19= 00FF0000      SA35= 001153DD      SA51= 00FF0000

SA 2= 004ACF82      SA18= 00FF0000      SA34= 00056B78      SA50= 00A06914

SA 1= 00CC295C      SA17= 00CC2521      SA33= 000926E8      SA49= 0007ADAB

SA 0= 00413A72      SA16= 003C2EBC      SA32= 0011A410      SA48= 00131768

DataFlash Checksum Total = 221AD4D2

Application Checksum     = 0933B869

BootLevel_1 Checksum     = 01A095D6

----------

TOTAL        = 2CEF2311

Intf Params checksum            = 00EF2838

STN-INTF(new) FPGA checksum     = 010DC3D0

CORR-INTF FPGA checksum         = 010CFE0A

STN-INTF(old) FPGA checksum     = 010EB4DA

SCC-CTRL FPGA checksum          = 011006DF

MEM-ADDR FPGA checksum          = 01040235

MEM-MUX FPGA checksum           = 00C1A903

MEM-DMUX FPGA checksum          = 010D63CE

----------------------------------------------

31.25 MHz TAPS checksum         = 00FED24B

62.5  MHz TAPS checksum         = 00FED149

----------------------------------------------

TFB-FIR-DLY FPGA checksum       = 010853DD

TFB-FIR-FIR  checksum           = 01BBE216

SCC:TFB-II:Quad0:Node16 > cc 32

ASYNC_ID 32 connected.

DTSSIM6:Quad0:Node32 > checksums

Usage:  checksum N:  N=1 for BOOTLEVEL_1,  =2 for APP,  =3 for SECTIONS,  = 4 for all.

Defaulting to ALL.

....

SECTION   START          LENGTH         CHECKSUM

-------   ----------     ----------     ----------

ICODE     0x00030000     0x00000146     0x00008B96

FCODE     0x000301AA     0x0002204C     0x012B0385

HCONST    0x000521F6     0x00018BA1     0x005EDB68

NCONST    0x0004C000     0x000000D4     0x000055EE

...BootLevel_1 chksum  = 0x01A095D6

...Application chksum  = 0x0B12E350

DTSSIM6:Quad0:Node32 > cc 33

ASYNC_ID 33 connected.

DTSSIM2:Quad0:Node33 > checksums

Usage:  checksum N:  N=1 for BOOTLEVEL_1,  =2 for APP,  =3 for SECTIONS,  = 4 for all.

Defaulting to ALL.

....

SECTION   START          LENGTH         CHECKSUM

-------   ----------     ----------     ----------

ICODE     0x00030000     0x00000146     0x00008B96

FCODE     0x000301AA     0x0002204C     0x012B0385

HCONST    0x000521F6     0x00018BA1     0x005EDB68

NCONST    0x0004C000     0x000000D4     0x000055EE

...BootLevel_1 chksum  = 0x01A095D6

...Application chksum  = 0x0B12E350

DTSSIM2:Quad0:Node33 > cc 36

ASYNC_ID 36 connected.

DPI:64ANT:Quad0:Node36 > fpga

SA15= 00FF0000      SA31= 00FF0000      SA47= 00FF0000      SA63= 00FF0000

SA14= 00FF0000      SA30= 00FF0000      SA46= 00FF0000      SA62= 00FF0000

SA13= 00FF0000      SA29= 00FF0000      SA45= 00FF0000      SA61= 00FF0000

SA12= 00FF0000      SA28= 00FF0000      SA44= 00FF0000      SA60= 00FF0000

SA11= 00FF0000      SA27= 00FF0000      SA43= 00FF0000      SA59= 00FF0000

SA10= 00FF0000      SA26= 00FF0000      SA42= 00FF0000      SA58= 00FF0000

SA 9= 00FF0000      SA25= 00FF0000      SA41= 00FF0000      SA57= 00FF0000

SA 8= 00FF0000      SA24= 00FF0000      SA40= 00FF0000      SA56= 00FF0000

SA 7= 00FF0000      SA23= 00FF0000      SA39= 00FF0000      SA55= 00FF0000

SA 6= 0075656C      SA22= 00FF0000      SA38= 00FF0000      SA54= 00FF0000

SA 5= 004A7374      SA21= 00FF0000      SA37= 00FF0000      SA53= 00FF0000

SA 4= 00600CDA      SA20= 00FF0000      SA36= 00FF0000      SA52= 00FF0000

SA 3= 00528D26      SA19= 00FF0000      SA35= 00C6AA7A      SA51= 00FF0000

SA 2= 0050B1C0      SA18= 00FF0000      SA34= 0009883E      SA50= 00FF0000

SA 1= 007561EC      SA17= 00FF0000      SA33= 00C6AE21      SA49= 00FF0000

SA 0= 0048E9F0      SA16= 00FF0000      SA32= 00093053      SA48= 00FF0000

DataFlash Checksum Total = 38EC81A8

Application Checksum     = 0BA80459

BootLevel_1 Checksum     = 01A095D6

----------

TOTAL        = 46351BD7

DPI-OUT FPGA checksum    = 00BFD8E0

DPI-FIFO FPGA checksum   = 01034BC0

DPI-IN FPGA checksum     = 00BE4BDC

DPI:64ANT:Quad0:Node36 > cc 35

ASYNC_ID 35 connected.

OPS:Quad0:Node35 > fpga

SA15= 00FF0000      SA31= 00FF0000      SA47= 00FF0000      SA63= 00FF0000

SA14= 00FF0000      SA30= 00FF0000      SA46= 00FF0000      SA62= 00FF0000

SA13= 00FF0000      SA29= 00FF0000      SA45= 00FF0000      SA61= 00FF0000

SA12= 00FF0000      SA28= 00FF0000      SA44= 00FF0000      SA60= 00FF0000

SA11= 00CC33CF      SA27= 00FF0000      SA43= 00FF0000      SA59= 00FF0000

SA10= 00447810      SA26= 00FF0000      SA42= 00FF0000      SA58= 00FF0000

SA 9= 00CC2454      SA25= 00FF0000      SA41= 00FF0000      SA57= 00FF0000

SA 8= 00447FCE      SA24= 00FF0000      SA40= 00FF0000      SA56= 00FF0000

SA 7= 00CC2B71      SA23= 00FF0000      SA39= 00FF0000      SA55= 00FF0000

SA 6= 00449748      SA22= 007689F0      SA38= 00FF0000      SA54= 00FF0000

SA 5= 00CC2146      SA21= 004AD734      SA37= 00FF0000      SA53= 00FF0000

SA 4= 00448673      SA20= 00760C35      SA36= 00FF0000      SA52= 00FF0000

SA 3= 0075F9CD      SA19= 004AEAC8      SA35= 00FF0000      SA51= 00FF0000

SA 2= 00488934      SA18= 00600CD3      SA34= 00FF0000      SA50= 00FF0000

SA 1= 00764E2F      SA17= 0052C7F3      SA33= 00FF0000      SA49= 00FF0000

SA 0= 004BAC29      SA16= 0050FB14      SA32= 00FF0000      SA48= 00FF0000

DataFlash Checksum Total = 351C5FC7

Application Checksum     = 0C3C30EE

BootLevel_1 Checksum     = 01A095D6

----------

TOTAL        = 42F9268B

FA-OUT-Y FPGA checksum    = 00C16124

FA-OUT-X FPGA checksum    = 00C0F6FD

FA-FIFO FPGA checksum     = 0103CFDA

BTypeA-Y FPGA checksum    = 0110ABDF

BTypeB-Y FPGA checksum    = 0110A422

BTypeB-X FPGA checksum    = 0110C2B9

BTypeA-X FPGA checksum    = 0110A7B9

FA-IN-Y FPGA checksum     = 00BE8301

FA-IN-X FPGA checksum     = 00C1FA58

OPS:Quad0:Node35 > cc 38

ASYNC_ID 38 connected.

PIC:64ANT:Quad0:Node38 > fpga

SA15= 00FF0000      SA31= 00FF0000      SA47= 00FF0000      SA63= 00FF0000

SA14= 00FF0000      SA30= 00FF0000      SA46= 00FF0000      SA62= 00FF0000

SA13= 00FF0000      SA29= 00FF0000      SA45= 00FF0000      SA61= 00FF0000

SA12= 00FF0000      SA28= 00FF0000      SA44= 00FF0000      SA60= 00FF0000

SA11= 00FF0000      SA27= 00FF0000      SA43= 00FF0000      SA59= 00FF0000

SA10= 00FF0000      SA26= 00FF0000      SA42= 00FF0000      SA58= 00FF0000

SA 9= 00FF0000      SA25= 00FF0000      SA41= 00FF0000      SA57= 00FF0000

SA 8= 00FF0000      SA24= 00FF0000      SA40= 00FF0000      SA56= 00FF0000

SA 7= 00FF0000      SA23= 00FF0000      SA39= 00FF0000      SA55= 00FF0000

SA 6= 00FF0000      SA22= 00FF0000      SA38= 00FF0000      SA54= 00FF0000

SA 5= 00FF0000      SA21= 00FF0000      SA37= 00FF0000      SA53= 00FF0000

SA 4= 00FF0000      SA20= 00FF0000      SA36= 00FF0000      SA52= 00FF0000

SA 3= 00FF0000      SA19= 00FF0000      SA35= 00FF0000      SA51= 00FF0000

SA 2= 00FF0000      SA18= 00FF0000      SA34= 00FF0000      SA50= 00FF0000

SA 1= 00FF0000      SA17= 00FF0000      SA33= 00FF0000      SA49= 00FF0000

SA 0= 00FF0000      SA16= 00FF0000      SA32= 00FF0000      SA48= 00FF0000

DataFlash Checksum Total = 3FC00000

Application Checksum     = 0C054F09

BootLevel_1 Checksum     = 01A095D6

----------

TOTAL        = 4D65E4DF

PIC-OUT FPGA checksum    = 01FE0000

PIC-FIFO FPGA checksum   = 02FD0000

PIC-IN FPGA checksum     = 01FE0000

PIC:64ANT:Quad0:Node38 > fpgaver

FPGA Version Number = 5

### Connectivity Robustness Tests

#### Correlator Card to Correlator Interface Card

Procedure:  This test verifies that the connections of the sum feedback to CAI-63 are functional and robust.  Use the “C 8” test available in the LTA. Use DPI commands to switch the input of antenna 63 on the CI 3 and 7 in each bin to the sum input.  This test verifies APP0190 and partly satisfies APP0200.

Results:  All inputs were verified.

Result of C 8 with CI switch in normal position.

Set all CI cards for identical 1 msec tics, normal data out (not PN).

CI#0  CI#1  CI#2  CI#3  CI#4  CI#5  CI#6  CI#7     SI#0  SI#1

----  ----  ----  ----  ----  ----  ----  ----     ----  ----

0 good:    0     0     0     0     0     0     0     0        0     0

1 good:   32    32    32    32    32    32    32    32      128   128

2 good:    0     0     0     0     0     0     0     0        0     0

3 good:    0     0     0     0     0     0     0     0        0     0

Node C00 CI Card# 0:     FPGA 0                                              FPGA 1

SWAP = 0                  SWAP = 1                   SWAP = 0                  SWAP = 1

-----------------------   -----------------------    -----------------------   ----------------------

0    90    180   270      0    90    180   270       0    90    180   270      0    90    180   270

IN  B0,B1 B0,B1 B0,B1 B0,B1   B0,B1 B0,B1 B0,B1 B0,B1    B0,B1 B0,B1 B0,B1 B0,B1   B0,B1 B0,B1 B0,B1 B0,B1

==  == == == == == == == ==   == == == == == == == ==    == == == == == == == ==   == == == == == == == ==

00  00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55    00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55

01  00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55    00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55

02  00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55    00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55

03  00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55    00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55

04  00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55    00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55

05  00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55    00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55

06  00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55    00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55

07  00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55    00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55

08  00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55    00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55

09  00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55    00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55

10  00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55    00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55

11  00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55    00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55

12  00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55    00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55

13  00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55    00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55

14  00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55    00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55

15  00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55    00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55

0 loops through all states, (      48 through all inputs)

Errors are CUMULATIVE.      c: s stop, 0-7 card, a tog  accum mode, d tog dsply mode, h tog hold, c clr

PHASE1 test active

The above verifies that the “normal” Correlator data path is running error-free.

Result of C 8 with CI switch in VLBI position (execute “setciswitch” at DPI node):

LTA:2Ant:Node0 >

CI#0  CI#1  CI#2  CI#3  CI#4  CI#5  CI#6  CI#7     SI#0  SI#1

----  ----  ----  ----  ----  ----  ----  ----     ----  ----

0 good:    0     0     0     2     0     0     0     2        2     2

1 good:   32    32    32    30    32    32    32    30      126   126

2 good:    0     0     0     0     0     0     0     0        0     0

3 good:    0     0     0     0     0     0     0     0        0     0

Node C00 CI Card# 3:     FPGA 0                                              FPGA 1

SWAP = 0                  SWAP = 1                   SWAP = 0                  SWAP = 1

-----------------------   -----------------------    -----------------------   ----------------------

0    90    180   270      0    90    180   270       0    90    180   270      0    90    180   270

IN  B0,B1 B0,B1 B0,B1 B0,B1   B0,B1 B0,B1 B0,B1 B0,B1    B0,B1 B0,B1 B0,B1 B0,B1   B0,B1 B0,B1 B0,B1 B0,B1

==  == == == == == == == ==   == == == == == == == ==    == == == == == == == ==   == == == == == == == ==

00  00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55    00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55

01  00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55    00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55

02  00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55    00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55

03  00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55    00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55

04  00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55    00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55

05  00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55    00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55

06  00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55    00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55

07  00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55    -- -- -- -- -- -- -- --   -- -- -- -- -- -- -- --

08  00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55    00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55

09  00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55    00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55

10  00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55    00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55

11  00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55    00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55

12  00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55    00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55

13  00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55    00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55

14  00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55    00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55

15  00,00 55,55 55,55 55,55   55,55 55,55 55,55 55,55    -- -- -- -- -- -- -- --   -- -- -- -- -- -- -- --

0 loops through all states, (     168 through all inputs)

Errors are CUMULATIVE.      c: s stop, 0-7 card, a tog  accum mode, d tog dsply mode, h tog hold, c clr

PHASE1 test active

Note that one card, CI-3, has only 30 good inputs whereas before it had 32.  Note in the above that two lines have dashes in FPGA-1.  These correspond to the CAI-63 inputs for both polarizations.  Since the input is open in this instance, dashes are shown by the test program.  This verifies that the switch for CAI-63 is working and thus verifies APP0190.

#### Correlator Card to PIC

Procedure:   Verify that the “sum” connections from the CC to the PIC are functional and robust.   From the Engineer Port, execute the following steps

- Set the CI data source to pseudo-random, non-repeating (cc 0, wrfd 1 0x3CF 0x400 19 0x80 1)

- At the PIC, execute pntest command and view the resulting error display for one minute.  There should zero errors in each channel, except for Bin 1, CC5, Rack 2 which uses TE as an input rather than the pseudo-random data.  This is because the 2-Antenna Correlator does not have enough available outputs to populate all PIC data inputs.

Results:

PIC:64ANT:Quad0:Node38 > pntest

Bin 0    Bin 0    Bin 1    Bin 1    Bin 2    Bin 2    Bin 3    Bin 3

CC1 CC1  CC5 CC5  CC1 CC1  CC5 CC5  CC1 CC1  CC5 CC5  CC1 CC1  CC5 CC5

Bit:      0   1    0   1    0   1    0   1    0   1    0   1    0   1    0   1

----------------------------------------------------------------------

Rack 1       0   0    0   0    0   0    0   0    0   0    0   0    0   0    0   0

Rack 2       0   0    0   0    0   0   99  99    0   0    0   0    0   0    0   0

Rack 3       0   0    0   0    0   0    0   0    0   0    0   0    0   0    0   0

Rack 4       0   0    0   0    0   0    0   0    0   0    0   0    0   0    0   0

The test passes. This is one of the tests that demonstrates compliance with APP0200

#### Ethernet to PICs

Procedure: Testing Ethernet connectivity to the PICs is done as a by-product of executing the temperature monitoring test described in Section 2.3.7.1.

Result:  The test was successful and demonstrates compliance with COR-0050.

#### TE from QCC to the PIC

Procedure:  This test verifies that the TE connection to the PICs from the QCC.

From the Engineering Port, execute the resette command at the PIC.  Then execute the seete command at the PIC.  See the example session below:

PIC:64ANT:Quad0:Node38 > resette

TE Error Flag has been cleared..

PIC:64ANT:Quad0:Node38 > seete

NO TE Error.

If each PIC shows “NO TE error” after an initial reset, the test passes.

Result:  The test passed.  There are no APP requirements directly related to this test but it is a pre-requisite to having the PICs correctly time tag the data.

#### Clock Cables for PIC.

Procedure: This procedure verifies that the 125 MHz clock connections to the PICs from the clock distribution network are functional.  These are implicitly tested in Section x.

Results: The test passed.  There are no APP requirements directly related to this test but it is a pre-requisite to having the PICs correctly time tag data.

### Environmental Tests

#### PIC Power Dissipation

Procedure:

ATX lead color codes and configuration are shown below.

Figure 2-12.  ATX power supply connector configuration.

Use the engineering port for configuring the 2-antenna correlator “Black Hole”.

Configure summers for counter output: ("cc 0" and "wrfd 1 0x3CF 0x400 19 0x20 1")

Configure PIC to transmit data: (“cc 36” and “startpic”)

Measure current on all voltage inputs using the connector configuration shown in Figure 2-12 as a guide.   Calculate individual powers and sum.

For the 48-volt bin voltage, the QCC “win” test to measure power with and without the PIC installed, and take the difference.

Assure that the “win” screen shows no out-of-tolerance voltages, currents or temperatures.

Results:

Equipment:

Meter used: General Technologies Corp. CM 100 current meter, S/N 41102397

Current Measurements:

| Voltage (VDC) | Current (amps) | Power (watts) |
| --- | --- | --- |
| 48 | 0.14 | 6.72 |
| 12 | 0 | 0 |
| 5 | 0.32 | 1.6 |
| +5 SB | 0.36 | 1.8 |
| 3.3 | 1.7 | 5.61 |
| -5 | 0 | 0 |
| -12 | 0 | 0 |
| Total |  | 15.73 |

Note: Above was done before the PIC personality used for AcRv was available.  The tests were repeated on 2014-11-3 with the final personality (= version 5) in place.  The results are summarized below

| Voltage (VDC) | Current (amps) | Power (watts) |
| --- | --- | --- |
| 48 |  | 3 |
| 12 | 1.32 | 15.84 |
| 5 | .72 | 3.6 |
| +5 SB | .357 | 1.785 |
| 3.3 | 1.73 | 5.709 |
| -5 | .03 | 0.15 |
| -12 | 0.012 | 0.144 |
| Total |  | 30.23 |

Win screen:

Figure 2-13.  The QCC "win 5" screen.  Since no voltages, currents or temperatures are in inverse video, they are all within tolerance.

From the above, we can extrapolate the effect of PIC power dissipation in the racks in which they are installed.  Typical Correlator rack power consumption, as measured at the Q4 Correlator PAS is 3700 watts.  Since the measured power of a PIC is about 30 watts, it is less than 1% of the rack power dissipation and is impossible to measure due to the noise in the current measurements.  The “win” screen from the 2-antenna Correlator is shown in Figure 2-13 shows no out-of-tolerance currents, voltages or temperatures (these would be highlighted in inverse video).  This test verifies the following requirements:

#### 1-PPS Power Dissipation

Procedure: The 1-PPS distributor is the only card in the bottom bin that is powered from a 9-U power card.  The 9-U card is located in the bin above it and the three required voltages, 1.8, 3.3 and 5.0 volts are provided via individual wires soldered from one backplane to the next.  To measure power then, simply measure the current in each of these three wires, similar to the method used above in Section 2.3.10.1 and compute the power.

Results:

The measurements and calculation are included in the table below:

This fulfills part of the following requirement: SIT-0020.

### Component Burn-in Tests

Components in the 64-Antenna Correlator have shown much a much lower than expected failure rate.  This is partly due to burn-in tests of those components in Charlottesville.  To help assure similar reliability, two types of components were burned in for the APP.  These include the ATX power supplies and the PICAs.  The details are summarized below.

#### ATX Power Supply Burn-in Tests

Procedure:  Load each output of each ATX supply at about 50% rated load and burn in for at least two days using a simple test fixture.  See Figure 2-14 below for a picture of the configuration.  Right-most is a 48-volt supply which is used to power the ATX supply.  The ATX supply is mounted on the chromated aluminum plated.  The test resistors used for loads are situated under the fans.  Check all voltages at the beginning and end of each test to assure that they are all within tolerance.

Figure 2-14.  Photo of ATX power supply burn-in test station.

Results: Results are summarized in the table below.

All power supplies passed the burn in test.  This partially fulfills the following requirement:APP0060

#### PICA Burn-in Tests

Procedure:  Burn-in each PICA in the Black Hole test fixture for at least 100 hours.  Perform the loop-back test described in RD-02 at the end of the burn-in to assure the unit survived the test.

Results:

All production PICA passed the burn-in test (serial numbers 1 and 2 were prototypes).  This partially fulfills the following requirement: APP0060.